Technical Specification of OT Light (High End)

| S.N. | Purchaser's Technical Specifications |  | Bidder's Compliance Sheet |  |  |
| --- | --- | --- | --- | --- | --- |
|  | OT Light (High End) |  | Yes/No | Page No. in Catalogue | Remarks |
|  | Manufacturer |  |  |  |  |
|  | Brand |  |  |  |  |
|  | Type/Model |  |  |  |  |
|  | Country of Origin |  |  |  |  |
| 1. | Description of Function |  |  |  |  |
|  | Surgical lights illuminate the surgical site for optimal visualization of small, low-contrast objects at varying depths in incisions and body cavities. |  |  |  |  |
| 2. | Operational Requirements |  |  |  |  |
|  | It shall be latest LED technology operating light field with 2 units, one major dome and one satellite dome. |  |  |  |  |
|  | Ceiling mounted LED Surgical Light with complete accessories. Operation Theatre Light, LED with all standard accessories. |  |  |  |  |
| 3. | Technical Specifications |  |  |  |  |
|  | The light shall adopt LED Technology to create a homogenous light patch without emitting any infrared rays. |  |  |  |  |
|  | Pulse width modulation control LED driving to ensure less heating of LED which increases life of LEDs and no change in Light color Output and light color temperature throughout life. |  |  |  |  |
|  | Lights should have higher watts LEDs to achieve high lumen to watt ratio which leads to lesser energy consumption and low heat at surgical area. |  |  |  |  |
|  | All Dome should have high CRI and get the required shade of light as per different surgical requirements. |  |  |  |  |
|  | Light intensity shall be adjustable between 25% -100%. |  |  |  |  |
|  | Both the dome must have low intensity light for endoscopy mode of 5% dimming or less. |  |  |  |  |
|  | Light Intensity, light field diameter and color temperature should be controlled from light arm control panel and remote-control panel or wall mount control. |  |  |  |  |
|  | Light head shall be of a shape to avoid laminar flow influence by turbulence. |  |  |  |  |
|  | The color temperature shall be synchronized by separate controlled touch panel integrated in the dome. |  |  |  |  |
|  | The light shall be mountable to celling from single center with 360 degree rotation of all arms. Spring arms shall be rotatable at least 360 degrees around its own axis. Each dome head should be rotatable with 360 degree at connecting joint with spring arm and at least 300 degree around its own axis. This feature should be applicable with camera mounted dome also in future upgradability. |  |  |  |  |
|  | The maximum movement angle of the spring arm shall be at least 45 degree down. |  |  |  |  |
|  | Each LED shall be replaceable individually to save cost in case of failure, instead of replacing the module with several LED’s. |  |  |  |  |
|  | Must have indicative alarm with failed module no. for easy replacement. |  |  |  |  |
|  | Future upgradability |  |  |  |  |
|  | Ready to use HD camera provision. |  |  |  |  |
|  | The detachable HD camera in ready with wireless technology for full HD recording and Live transmission should be available in future provision. |  |  |  |  |
|  | Should have future provision of two-way audio transmission with video conferencing. |  |  |  |  |
|  | Should have obstacle sensing to maintain the intensity without temperature rise. |  |  |  |  |
|  | Major Dome |  |  |  |  |
|  | Should have 2 color LED lamps, White & Amber with endoscopic mode. |  |  |  |  |
|  | Central Illuminance should be 160,000 lux |  |  |  |  |
|  | Light field Diameter should be adjustable form 150mm to 300mm. |  |  |  |  |
|  | Color temperature (K), adjustable 3500 – 5000K. |  |  |  |  |
|  | The focal depth should be from 650 mm to 1200 mm. |  |  |  |  |
|  | Color rendering index Ra should be 93 or more and R9 should be 92 or more. |  |  |  |  |
|  | Endoscopy mode illumination should be less than 5% of 160,000 lux. |  |  |  |  |
|  | Number of LED bulbs the light head should not be more than 48. |  |  |  |  |
|  | LED Service life should be more than 60000hrs. |  |  |  |  |
|  | Max. power consumption of all light sources should be less than 100W at maximum setting. |  |  |  |  |
|  | Should have centralized control system and electronic focusing. |  |  |  |  |
|  | The lamp head diameter must be 590mm or more. |  |  |  |  |
|  | Should provide sterilizable and detachable light control handle. |  |  |  |  |
|  | The radiant energy must be less than 3.2 mW/lm2 so that life of the LED increased. |  |  |  |  |
|  | The temperature rise at the surgeon head level should be less than 20C. |  |  |  |  |
|  | It must have a quality LED light and light color outputs should not be changed throughout life cycle. |  |  |  |  |
|  | Should supply either wall mount control or Remote control for each dome. |  |  |  |  |
|  | Satellite Dome |  |  |  |  |
|  | Should have single color LED lamps with endoscopic mode. |  |  |  |  |
|  | Central Illuminance should be 160,000 lux |  |  |  |  |
|  | Light field Diameter should be adjustable form 150mm to 300mm. |  |  |  |  |
|  | Color temperature (K): 4000K. |  |  |  |  |
|  | The focal depth should be from 590 mm to 1000 mm. |  |  |  |  |
|  | Color rendering index Ra should be 95 or more and R9 should be 93 or more. |  |  |  |  |
|  | Endoscopy mode illumination should be less than 5% of 160,000 lux. |  |  |  |  |
|  | Number of LED bulbs the light head should not be more than 24. |  |  |  |  |
|  | LED Service life should be more than 60000hrs. |  |  |  |  |
|  | Max. power consumption of all light sources should be less than 80W at maximum setting. |  |  |  |  |
|  | Should have centralized control system and electronic focusing. |  |  |  |  |
|  | The lamp head diameter must be 680mm or more. |  |  |  |  |
|  | Should provide sterilizable and detachable light control handle. |  |  |  |  |
|  | The radiant energy must be less than 3.2 mW/lm2 so that life of the LED increased. |  |  |  |  |
|  | The temperature rise at the surgeon head level should be less than 20C. |  |  |  |  |
|  | It must have a quality LED light and light color outputs should not be changed throughout life cycle. |  |  |  |  |
|  | Should supply either wall mount control or Remote control for each dome. |  |  |  |  |
| 4 | Accessories, Spare Parts and Consumables included |  |  |  |  |
|  | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above). |  |  |  |  |
| 5 | Operational Environment |  |  |  |  |
|  | The product offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |  |  |  |  |
|  | Power Supply: 220 – 240V AC, 50 Hz with appropriate power cable and plug. |  |  |  |  |
|  | The length of the power cable must be at least 3 m. |  |  |  |  |
| 6. | Standards, Medical Device Regulation and Safety Requirements |  |  |  |  |
|  | Must have product compliance of IEC 60601-1, IEC 60601-2, IEC 60601-2-41 |  |  |  |  |
|  | Must submit a valid "Medical devices – Quality management systems – Requirements for regulatory purpose"; EN ISO 13485:2016. |  |  |  |  |
|  | Must submit valid European CE certificate (Directive 93/42 EEC on Medical Devices) and US-FDA resisted product. |  |  |  |  |
| 7. | User and Technical Training |  |  |  |  |
|  | Must provide user training (including how to use and safe handling, keep hygiene the equipment). |  |  |  |  |
|  | Must provide the first step technical training to troubleshooting and maintenance of the machine to the hospital technician and engineers. |  |  |  |  |
| 8. | Up time Service Backup, Warranty/Guarantee |  |  |  |  |
|  | Should have 1 year warranty and 1 year free service. |  |  |  |  |
|  | During the warranty period supplier must ensure preventive maintenance and corrective/breakdown maintenance whenever required. |  |  |  |  |
| 9. | Installation and commissioning |  |  |  |  |
|  | The bidder must arrange for the quoted equipment with all the listed accessories and consumables to be installed on site and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |  |  |  |  |
| 10. | Documentation on site |  |  |  |  |
|  | User (Operating) manual in English |  |  |  |  |
|  | Service (Technical/ Maintenance) manual in English |  |  |  |  |
|  | Spare parts and consumables price list should be provided in the separate sheet, non-submission leads to the rejection of the bid. |  |  |  |  |