Technical Specification of Electro-hydraulic OT Table (High End)

| S.N. | Purchaser's Technical Specifications |  | Bidder's Compliance Sheet |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Electro-hydraulic OT table |  | Yes/No. | Page No. in Catalogue | Remarks |
|  | Manufacturer |  |  |  |  |
|  | Brand |  |  |  |  |
|  | Type/Model |  |  |  |  |
|  | Country of Origin |  |  |  |  |
| 1. | Description of Function |  |  |  |  |
|  | Fully electro-hydraulic operating tables are tables for performing surgical procedures and it works with electrical power and hydraulic system. |  |  |  |  |
| 2. | Operational Requirements |  |  |  |  |
|  | OT Table is required for general surgery, Laparoscopy,bariatric surgery, CTVS, Spinal surgery, Neurosurgery, ENT, Urology, Ob/Gyno and shall have X-Ray translucent tops. |  |  |  |  |
|  | Electric and hydraulic Operating Table with complete standard accessories. |  |  |  |  |
| 3. | Technical Specifications |  |  |  |  |
|  | Five section table topincluding divided foot section. |  |  |  |  |
|  | It should made of SS-304 medical Class materials with Durable and Corrosion resistant properties. |  |  |  |  |
|  | The form pad should be memory form with 60 mm thickness with seamless PU leather. |  |  |  |  |
|  | The form should have antibacterial, antistatic, and waterproof properties. |  |  |  |  |
|  | The lifting column should be protected by elastic plastic cover that helps for easy cleaning and aesthetically looking. |  |  |  |  |
|  | The hydraulic system should be high pressure with integrated valve and 24V DC motor. |  |  |  |  |
|  | Table wheel should be inside the Base frame |  |  |  |  |
|  | Table should have electronic Floor Locking system for Stability and movements through lock/unlock floor feet button. |  |  |  |  |
|  | Table should have Pneumatic push Button for Leg plate assembly. |  |  |  |  |
|  | Leg Plate should be controlled by Gas/pneumatic Spring operated. |  |  |  |  |
|  | The keyboard must have a button for zero position. |  |  |  |  |
|  | The keyboard must have lock and unlock button to prevent the accidental press to the button. |  |  |  |  |
|  | Table should have Dual Control Panel of conventional handheld control and remote control. |  |  |  |  |
|  | Table Should be C-Arm Compatible |  |  |  |  |
|  | The Table must have easy click leg plate control by electro-hydraulic and manual. |  |  |  |  |
|  | The table must have beach chair position. |  |  |  |  |
|  | Tabletop must have full length radio translucent for free C-Arm access made of phenolic laminate. |  |  |  |  |
|  | All table positioning, i.e., height, back section, lateral tilt, Trendelenburg, and reverse-Trendelenburg, height adjustment except foot and head section must be operated electrically. |  |  |  |  |
|  | It should have maximum lifting capacity of 300 kg. |  |  |  |  |
|  | The table top sliding (electrical longitudinal shift function) must be minimum of 300 mm for free access to C-Arm |  |  |  |  |
|  | There shall be the provision of attachment of the fracture table (Optional). |  |  |  |  |
|  | Dimensions (approx.) |  |  |  |  |
|  | Lowest Height: 700±20mm |  |  |  |  |
|  | Height adjustment: 350 mm |  |  |  |  |
|  | Left/Right tilt: 200/200 |  |  |  |  |
|  | Trendelenburg: 30º |  |  |  |  |
|  | Reverse -Trendelenburg: 30º |  |  |  |  |
|  | Leg upward/downward: 900/900 |  |  |  |  |
|  | Back boar upward/Downward: 750/350 |  |  |  |  |
|  | Maximum width: 520 mm or more |  |  |  |  |
|  | Length: 2050 mm or more |  |  |  |  |
|  | Hand board upwards: 460 or more |  |  |  |  |
|  | Mechanical Function |  |  |  |  |
|  | Kidney elevator: 120 mm |  |  |  |  |
|  | Head board: detachable |  |  |  |  |
|  | Leg board: detachable |  |  |  |  |
| 4 | Accessories, Spare Parts and Consumables included |  |  |  |  |
|  | All standard accessories, consumables and parts required to operate the equipment should be supplied. |  |  |  |  |
|  | Hand remote -1 |  |  |  |  |
|  | Foot support -1 pair |  |  |  |  |
|  | Arm support – 1 pair |  |  |  |  |
|  | Body strap -1 |  |  |  |  |
|  | Shoulder support – 1 pair |  |  |  |  |
|  | Leg support -1 pair |  |  |  |  |
|  | Body support -1 |  |  |  |  |
|  | Fixing clamp -1 |  |  |  |  |
|  | Mattress – 1 |  |  |  |  |
|  | Anesthetic screen -1 |  |  |  |  |
| 5 | Operational Environment |  |  |  |  |
|  | The product offered shall be designed to be stored and to operate normally under the conditions of the purchaser's country. The conditions include Power Supply, Climate, Temperature, Humidity, etc. |  |  |  |  |
|  | Power Supply: 220 ±22 VAC, 50 Hz with appropriate power cable and plug. |  |  |  |  |
|  | The internal battery backup must be at least 50-80 operations or about one-week normal operations in case of external electricity failure. |  |  |  |  |
|  | The length of the power cable must be at least 2-3 m. |  |  |  |  |
| 6. | Standards, Medical Device Regulation and Safety Requirements |  |  |  |  |
|  | Must have product compliance of IEC 60601-1, IEC 60601-2, IEC 60601-2-41 |  |  |  |  |
|  | Must submit a valid "Medical devices – Quality management systems – Requirements for regulatory purpose"; EN ISO 13485:2016. |  |  |  |  |
|  | Must submit valid CE certificate (Directive 93/42 EEC on Medical Devices) and/or US-FDA approved. |  |  |  |  |
| 7. | User and Technical Training |  |  |  |  |
|  | Must provide user training (including how to use and safe handling, keep hygiene the equipment). |  |  |  |  |
|  | Must provide the first step technical training to troubleshooting and maintenance of the machine to the hospital technician and engineers. |  |  |  |  |
| 8. | Up time Service Backup, Warranty/Guarantee |  |  |  |  |
|  | Should have 1 year replacement of parts & service warranty and another 1 year of free service warranty from the date of complete installation. (Delivery and installation of machine of all the item as per tender.) |  |  |  |  |
|  | During the warranty period supplier must ensure preventive maintenance and corrective/breakdown maintenance whenever required. |  |  |  |  |
| 9. | Installation and commissioning |  |  |  |  |
|  | The bidder must arrange for the quoted equipment with all the listed accessories and consumables to be installed on site and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail. |  |  |  |  |
| 10. | Documentation on site |  |  |  |  |
|  | User (Operating) manual in English |  |  |  |  |
|  | Service (Technical/ Maintenance) manual in English |  |  |  |  |
|  | Spare parts and consumables price list should be provided in the separate sheet, non-submission leads to the rejection of the bid. |  |  |  |  |